Eddy Covariance Fieldwork Checklist

Technician Name: _____________________________________

Site ID: _____________________________________

Arrival Date/Time (UTC-7, Phoenix, AZ time): ________________________________________

Every visit (correct and note any issues or actions taken)

General checks

- Photograph the tower and any issues

- Check for any obstructions accessing the tower (loose debris, wires, tubing)

- Check for structural damage to the tower or instruments

- Check for new animal activity (rodent burrows, wasps, etc)

- Check water or insect intrusion into sensors + enclosures

- Clear snow/debris on solar panels and other instruments, esp. radiometers

Data acquisition system

- Check power system voltage:  ______________________

- Check datalogger clocks

Datalogger date/time: ______________________________

Actual date/time (UTC-7, Phoenix, AZ time): _____________________________

Adjust clock if >2min off

- Check for datalogger error messages

- Swap out SD card

  - Press “Eject,” wait for “Act” light to turn green

  - Remove SD card, store somewhere safe

  - Insert fresh SD card, wait for “Act” light to start slowly flashing red

  - “Act” light does not turn yellow/orange

Leaving

- Enclosures are closed + locked

- If enabled, ensure remote communication is functioning

Departure Date/Time (UTC-7, Phoenix, AZ time): __________________________________

Notes

Notes (cont.)

Eddy Covariance Monthly + Semi-Annual Preventative Maintenance Checklist

Technician Name: _____________________________________

Site ID: _____________________________________

Monthly Preventative Maintenance (correct and note any issues or actions taken)

- Check cables/conduits for damage

- Check the sonic anemometer is level and facing the proper direction

- Check the radiometers are level

- Check for heavily trampled vegetation under net radiometer

Semi-Annual Preventative Maintenance

- Inspect and test batteries

- Gently clean the surfaces of the gas analyzer and radiometers with an alcohol wipe

- Check if any sensors need to be calibrated following calibration interval recommendations

- Remove broken/malfunctioning/deprecated sensors

- Make sure all serial numbers and firmware versions are documented

- Verify the orientation, separation, and measurement heights of EC instruments

Notes

Notes (cont.)

Eddy Covariance Raw Data QA/QC Checklist

Technician Name: ______________________________

Data download date: ______________________________

Today’s date: _________________________

Site ID: _________________________

*** Note: this is a very rough guide to help flag immediate problems with raw data for further scrutiny. Data quality issues can persist even if all checks pass, and data that fails these checks may actually be OK.

General checks (note any failed checks)

- Data record starts when expected (SD card install date)

- Battery (batt)

  - never falls below 11.5V, never exceeds 14.5V

  - Reaches full recharge each day (doesn’t slowly decline/weaken over the course of the record)

- Eddy Covariance:

  - disable_sonic ~ 0 for all datapoints, is not consistently > 3600

  - disable_irga ~ 0 for all datapoints, is not consistently > 3600

  - rho_co2 between 500 – 800 mg/m3

  - rho_h2o < 16 g/m3 in winter, <40 in summer, never < 0

  - irgason_ws within reasonable values (Avg < 10m/s, Max < 65 m/s)

  - irgason_wd within reasonable values (always between 0 and 360)

  - Occasional large spikes are OK, but persistent bad data/non-physical values are a concern

- Temperature variables (panel_temp, sonic_temp, irgason_temp, cnr4_temp, temp, si111_temp, surf_temp, soil_temp)

  - Variables do not deviate from each other over the course of the collection period

  - Variables follow expected diurnal cycle (lowest before dawn, highest in the early afternoon)

  - Values are within reasonable meteorological ranges (15-35C in Summer, -20 – 10C in Winter)

  - No long runs of NA values

- Radiation variables

  - sw_out < sw_in

  - sw_out, sw_in, ppfd ~ 0 at night

  - sw_in peaks around 1000 W/m2 (Summer) or 500 W/m2 (Winter) during the day

  - ppfd peaks around 2000 umol/m2/s (Summer) or 1000 umol/m2/s (Winter) during the day

  - lw_in, lw_out follow a diurnal cycle, higher in the daytime, around 230 W/m2 in winter and 450 W/m2 in the summer

- Soil variables

  - vwc ranges from 0.05 to 0.5

  - vwc spikes in response to precipitation across all pits, with a delayed response in deeper sensors

  - G is near zero or barely negative in the winter, and follows a diurnal cycle in the summer (positive, <20W/m2 in the daytime)

- Other variables

  - 75kPa < irgason_press, press < 85kPa

  - 0 < vapor_press < ~3kPa

  - 0 <  rh < 100%

- SD Card quick formatted after copying data over to local machine (FAT32, 16kb allocation size)

Notes